CTA GI Bleed 01/13/2025

CTDI:  <25 mGy per acquisition

PT Preparation:

Setup:   Supine, AP Scout from above the diaphragm through the ischial tuberosities

DFOV:   Appropriate for patient’s body habitus

Scan Parameters:

Non-Contrast

- Scan from above the diaphragm through the ischial tuberosities

Late- Arterial Phase

- Scan from above the diaphragm through the ischial tuberosities

- IV Contrast at the discretion of the Radiologist

  - 1.5 mL/kg body weight of 320-370 mg iodine /ml non-ionic contrast (Not to exceed 150ml) @ a rate of 4 ml/sec

  - Followed by up to 50 ml normal saline @ a rate of 4 ml/sec

- In order to obtain imaging in the appropriate phase please follow the steps below

  - First step is to set a pre-defined bolus-trigger of 150 HU

  - Second step is to place an ROI in the abdominal aorta and add a delay of 12 seconds

    - The goal is to start the scan 12 seconds after the ROI has reached 150 HU

Delay ~ 150 seconds post arterial phase

- Scan from above the diaphragm through the ischial tuberosities

PACS Series:

- Topogram

- 5x5 Non-Contrast Axial

  - 1.25x1.25 Non-Contrast Coronal MPR

- 1.25 x 1.25 Arterial Axial

  - 1.25 x 1.25 Arterial Coronal MPR

  - 1.25 x 1.25 Arterial Sagittal MPR

  - 10 x 1.25 Arterial Thin MIP Coronal

  - 10 x 1.25 Arterial Thin MIP Sagittal

- 5x5 Delay Axial

  - 1.25x1.25 Delay Coronal MPR